## Slide 1: Advanced Institute on Health Investigation and Air Sensing for Asian Pollution (AI on Hi-ASAP)
Chun-Hu Liu 2020/10/05

## Slide 2: Outline
Get calibration factor from reference instrument
  Rational of calefaction factor
  How to run python code to set calibration factor
Data cleaning
  What data cleaning tasks need to be done
  How to run python code to do data cleaning

## Slide 3: Get calibration factor from reference instrument

## Slide 4: Rationale of calibration factor
1.AS-Lung compare with reference instrument
Reference instrument
AS-Lung
Rename the filename of AS-Lung
Ex: 2020-09-27.csv AL-0001_2020-09-27.csv
The filename should be contain AS-Lung ID
(AL-0001 or AL0001)
Rename the filename of reference standard
Ex: Grimm026.xls standard_Grimm026.xls
The filename should be contain “standard”
and with excel data file
Data format of AS-Lung
Do not modify the data format of AL-Lung
Data format of reference instrument
Rename the column name of time and PM
Modify the data time format of time column
2.Rename the filename to fit the python code
3.Rename the column name to fit the python code
IMPORTANT
Column name and data time format of should be the same as the figure

## Slide 5: Rationale of calibration factor
AS-Lung compare with reference instrument
Reference instrument
AS-Lung
Rename the filename of AS-Lung
Ex: 2020-09-27.csv AL-0001_2020-09-27.csv
The filename should be contain AS-Lung ID
(AL-0001 or AL0001)
Rename the filename of reference standard
Ex: Grimm026.xls standard_Grimm026.xls
The filename should be contain “standard”
and with excel data file
Data format of AS-Lung
Do not modify the data format of AL-Lung
Data format of reference instrument
Rename the column name of time and PM
Modify the data time format of time column
Rename the filename to fit the python code
Rename the column name to fit the
python code
Select regression model
Simple linear regression or Two segments regression
Calibration factor
Factors of PM1 and PM2.5 will be reported since PM10 of G3 sensor is not reliable.

## Slide 6: Regression model
Simple linear regression
Two segments regression
Breaking point
When run the python code, we can select regression model.

## Slide 7: How to run the Python Code (1)
1. Open PyCharm -> project -> run.py
Step 1
Step 2
Step 3
Step 1:
Copy raw data to the folder of “ASLung_Calibration_Factor_RawData”
Step 2:
Open run.py (double click)
Step 3:
Run the python code

## Slide 8: How to run the Python Code (2)
2. Select 1) ASLung calibration factor from reference instrument to run the python code
3. Select Regression model and input PM value after “Maximum concentration of standard PM”
4. The python code will automatically average data to 1 minutes.
5. The python code also select the calibration range from 1 to max value of input after the maximum value of the data set
3.1 Maximum concentration of standard PM is the highest concentration of the regression.
3.2 We use 200 (chamber) or 150 (hood) as maximum value.
3.2 If max value is 200, the calibration range is from 1 to 200.
Notes: The select window is below the python code window

## Slide 9: How to run the Python Code (2)
Notes: The select window is below the python code window

## Slide 10: How to run the Python Code (3)
6. When finish, the file will be save in the folder of “ASLung_Calibration_Factor “
7. Open the file, copy the calibration factor and past them to the google sheet.
https://docs.google.com/spreadsheets/d/1yuvjPvsr1sEzm_pXpZWpnMxEC3WbsKn65V-apohp86E/edit#gid=0
This google sheet only for demo and course, all the research group has individual google sheet
The format between csv file and google sheet are the same, DONOT change the data format
Notes: The select window is below the python code window

## Slide 11: Google sheet
If you renew the calibration factor, please key in the end date in the “End_date” column.
Add a new row to add new calibration factor, DO NOT replace the old factor

## Slide 12: Data cleaning

## Slide 13: What data cleaning tasks need to be done
AS-Lung monitoring data
There are two ways toe get raw data
1.data in SD card or 2.Data from database
Data cleaning criteria
Step of data cleaning
PM < 1
PM > 50 and PM1=PM2.5=PM10
Remove ghost peak
Before
After
Set PM value as NaN
NaN is equal null

## Slide 14: Remove ghost peak
Average PM (AVGPM) of the 5 data before or after the ghost peak
Ghost peak (G)
10

## Slide 15: What data cleaning tasks need to be done
Before
After
4. Remove data of temperature, humidity
and CO2 when values are less than 1
5. Get calibration factor from google drive
and calibrate AS-Lung data
6. If calibrated PM1 > PM2.5, PM1 = PM2.5
7. If the missing data is more than 1/3 in an
hour, the python code will auto remove all
the data in the hour
Calibration factor in google sheet
Calibration factor API from google sheet
IMPORTANT
Do not modified the format of the google sheet

## Slide 16: Google sheet

## Slide 17: Calibration factor API
https://script.google.com/macros/s/AKfycbwfhUbpNqk5AE4HpUg0Dp-0pT1oMKa1mxLzWWAXb3dlnhTYRN8/exec
We create calibration factor API from google sheet. Keep the format of the google sheet is very important.

## Slide 18: Data prepare
1.Folder name of the data set should be contain AS-Lung ID, EX:AL-0001
2.Rename the filename is optional
AS-Lung ID only
Or
Lab ID before AS-Lung ID , separate by “_”
Keep the original filename
Or
Rename the filename
it is optional

## Slide 19: How to run data cleaning
3. Open PyCharm -> project -> run.py
Step 1
Step 3
Step 2
Step 1:
Copy raw data to the folder of “ASLung_RawData”
Step 2:
Open run.py (double click)
Step 3:
Run the python code

## Slide 20: How to run data cleaning
4. Select 2) Data cleaning and calibrate ASLung raw data
5. Select your data source, SD card or database
6 . Calibrated data will be save in the folder of
“ASLung_Calibrated_Data”, the filename will add prefix of “cal_”
  in the front of original filename

## Slide 21: Step of data cleaning will remind again when run the python code

## Slide 22: Open CSV file after calibrated
Open the calibrated csv file, AS-Lung ID, lab id , calibrated PM1 and PM2.5 will add in the end
column of the data set

## Slide 23: How to update calibration factor API link
1.Open run.py
2.Go to line 5, find calfactorurl = “xxxxxxxxxxx”
3. Past the new link after calfactorurl=“
You can get your google sheet and calibration factor API link from “GoogleSheetAndAPI.xlsx

## Slide 24
Thanks!